České vysoké učení technické v Praze

Fakulta jaderná a fyzikálně inženýrská

Czech Technical University in Prague

Faculty of Nuclear Sciences and Physical Engineering

doc. Ing. Jan Kučera, CSc.

Úloha neutronové aktivační analýzy

v chemické metrologii

The role of neutron activation analysis

in chemical metrology

Summary

Neutron activation analysis (NAA) plays an indispensable role in chemical metrology, namely in the preparation of reference materials (RMs) of chemical composition and in quality assurance due to many advantages, which are briefly reviewed. Recently, NAA has been recognized as a primary method of measurement. Hence, it can be expected that the role of NAA in chemical metrology will become even more important.

A special feature of NAA to verify analytical data that the technique produces has been termed the self-verification principle. The strategies of the self-verification principle are explained and examples of its utilization are presented.

One of the requirements for a primary method of measurement is that a complete uncertainty statement can be written down in terms of SI units (or ratios). The author was one of the first specialists, who dealt with uncertainty quantification in NAA and his approach to this topic is presented in this work.

Activities in utilization of NAA in the preparation of RMs, i.e., homogeneity testing and certification analyses of element contents, are listed and several important examples are presented. The RMs tested and analyzed were produced by national bodies, as well as those from abroad, including the most important world-leading institutions, namely the US National Institute of Standards and Technology (NIST) and the Institute of Reference Materials and Measurements (IRMM – Joint Research Centre of the European Commission).

Finally, use of NAA in the detection of errors is highlighted and several examples of such application of NAA are presented. One of them concerns revealing biased values for Mn and V in NIST Standard Reference Material (SRM) 1648 Urban Particulate Matter. New certified values for these elements were established by taking into account the author´s results. This was only the second case, where data for environmental and/or biological RMs obtained outside NIST has led to a correction of the original NIST certified values of element contents.

Souhrn

Neutronová aktivační analýza (NAA) má nezastupitelnou úlohu v chemické metrologii, zejména při přípravě referenčních materiálů (RM) chemického složení a zabezpečení jakosti pro mnoho svých výhod, které jsou stručně shrnuty. Nedávno byl NAA přiznán statut primární měřicí metody. Lze tedy očekávat, že úloha NAA v chemické metrologii se stane ještě významnější.

Speciální schopnost NAA verifikovat analytická data, která tato technika produkuje, byla nazvána samoověřovacím principem. Je vysvětlena strategie samoověřovacího principu a jsou uvedeny příklady jeho využití.

Jedním z požadavků na primární měřicí metodu možnost podat kompletní výčet zdrojů nejistot v jednotkách SI nebo v jejich poměrech. Autor byl jedním z prvních odborníků, kteří se zabývali kvantifikací nejistot v NAA. Jeho poznatky v této problematice jsou uvedeny v této práci.

Jsou shrnuty aktivity při využití NAA při přípravě RM, tj. při testování homogenity a certifikačních analýzách obsahů prvků a jsou uvedeny některé důležité příklady. Testované a analyzované RM byly vyrobeny jak v ČR (dříve v Československu), tak v zahraničních institucích, včetně nejvýznamnějších světových producentů, jimiž jsou americký National Institute of Standards and Technology (NIST) a Institute of Reference Materials and Measurements (IRMM – Joint Research Centre of the European Commission).

Závěrem je poukázáno na použití NAA při odhalování chybných analytických dat a je uvedeno několik příkladů takového využití NAA. Jeden z nich se týká odhalení chybně stanovených hodnot pro Mn a V  v  NIST standardním referenčním materiálu (SRM) 1648 Urban Particulate Matter. Nově certifikované hodnoty byly stanoveny s přihlédnutím k autorovým výsledkům. Jedná se teprve o druhý případ, kdy data pro environmentální nebo biologické RM získaná mimo NIST vedla ke korekci původních hodnot certifikovaných v NIST.

Klíčová slova

Neutronová aktivační analýza, chemická metrologie, primární měřicí metoda, princip samoověřování, nejistota, referenční materiál chemického složení, testování homogenity, certifikační analýza

Key words

Neutron activation analysis, chemical metrology, primary method of measure-ment, self-verification principle, uncertainty, reference material of chemical composition, homogeneity testing, certification analysis

©, ISBN

CONTENTS

1	INTRODUCTION									 6

1.1	NAA as a primary method of measurement				 7

1.2 	Use of NAA in the preparation of reference materials		10

2	EXPERIMENTAL								12

3	RESULTS AND DISCUSSION						14

3.1	Self-verification principle of NAA					14

3.2	Uncertainty quantification in NAA					16

3.3	Use of NAA in the certification of reference materials		19

3.4	Use of NAA in the detection of errors					25

4	CONCLUSIONS AND OUTLINE OF FUTURE RESEARCH

AND EDUCATIONAL ACTIVITIES					27

Acknowledgements									29

References											30

Curriculum vitae										35

1	INTRODUCTION

Neutron activation analysis (NAA) is based on the measurement of characteristic radiation, most frequently gamma-rays, of radionuclides formed by irradiation of the material with neutrons. The highest sensitivity for most elements is obtained when irradiation is carried out in the most powerful neutron source – in a nuclear reactor. Due to a higher potential for accuracy compared with other methods of elemental analysis, especially for trace element analysis, NAA has been denoted for a long time “reference, arbitrary, independent” method due to the following advantageous features of the technique [1]:

- sensitivity and applicability for minor and trace elements in a wide range of matrices;

- the virtual absence of an analytical blank;

- the relative freedom from matrix and interference effects;

- the possibility to perform analysis nondestructively using so-called instrumental neutron activation analysis (INAA);

- high specificity based on the individual characteristics of the induced radionuclides, such as (-ray energies and half-lives;

- the capability of INAA for multi-element determination, often allowing 30 to 40 elements to be determined in many matrices;

- the possibility to analyze the sample bulk due to the high penetrating power of neutrons and emitted (-ray, which is of special importance for sample difficult to dissolve;

- an inherent potential for accuracy compared to other analytical techniques; since the theoretical basis of INAA is simple and well understood, the sources of uncertainty can be modelled and well estimated;

- the totally independent principle as a nuclear-based property in contrast to the electronic nature of most other analytical techniques;

- the isotopic basis which often offers a choice of analytically independent routes for element determination - the self-verification principle [2].

In cases where the induced radionuclides of trace elements are masked by matrix activity, radiochemical separation provides interference-free detection limits close to the theoretical ones. Thus, in the radiochemical mode of NAA (RNAA), the technique has other advantageous features:

- trace and ultratrace (radio) chemistry can be performed under controlled conditions by inactive carrier additions;

- the chemical yield of separation can be obtained simply using carrier budgeting or the radiotracer method.

For these reasons, NAA has been used extensively used in chemical metrology, namely in quality assurance, in the preparation of reference materials, and for the detection of errors [3]. However, only recently [4] the Consultative Committee on the Amount of Substance – Metrology in Chemistry (Comité consultatif pour la quantité de matière – métrologie en chemie – CCQM) has agreed to consider NAA (using the comparator method of standardization) as a primary method of measurement.

1.1		NAA as a primary method of measurement

Fig. 1 shows the hierarchy of international metrological bodies. The Metre Convention is a treaty that created the International Bureau of Weights and Measures (BIPM), an intergovernmental organization under the authority of the General Conference on Weight and measures (CGPM) and the supervision of the International Committee for Weights and Measures (CIPM). The BIPM acts in matters of world metrology, particularly concerning the demand for measure-ment standards, and the need to demonstrate equivalence between national measurement standards.

Fig. 1. Hierarchy of metrological bodies	Fig. 2. List of CIPM Consultative      Committees

The CIPM has set up a number of Consultative Committees to consider advances that directly influence metrology (Cf. Fig. 2). For metrology in chemistry, the Consultative Committee on the Amount of Substance (CCQM) has been established in 1993.

CCQM defines a primary method of measurement as follows[5,6]:

“A primary method of measurement is a method having the highest metrological properties, whose operation can be completely described and understood, for which a complete uncertainty statement can be written down in terms of SI units.

- A primary direct method: measures the value of an unknown without reference to a standard of the same quantity.

- A primary ratio method: measures the value of a ratio of an unknown to a standard of the same quantity; its operation must be completely described by a measurement equation.”

Until 2008, the five methods designated by CCQM as potentially primary methods of measurement included: isotope dilution mass spectrometry (IDMS), coulometry, gravimetry, titrimetry and determination of freezing point depression [5]. Of the five methods listed, only IDMS, as ID-inductively coupled plasma mass spectrometry (ID-ICPMS), is capable of determination of trace elements in complex matrices, as is NAA. The case for NAA as a primary method has been made to CCQM by Bode, Fernandes and Greenberg [4] since 2000. Although the final recommendations are not yet finalized, it appears that CCQM has agreed to consider NAA as a primary method of measurement [4].

The consideration of NAA as a primary method of measurement has been achieved in the following way. Rearranging the above definition, it appears that a primary method of measurement must meet three criteria:

- The method’s operation can be completely understood, and fully described by a measurement equation;

- A complete uncertainty statement can be written down in terms of SI units (or ratios);

- The method has the highest metrological properties.

Meeting the first two requirements appears relatively straightforward for NAA using the comparator (relative) method of standardization. Of many various ways, the NAA measurement process can be most conveniently described for the purpose by the following equation:

(1)

where = mass of an element in the unknown sample; = mass of an element in the comparator standard; = ratio of isotopic abundances for unknown and standard; = ratio of neutron fluences (including fluence drop off, self-shielding, and scattering); = ratio of effective cross sections if neutron spectrum shape differs from unknown to standard; = ratio of counting efficiencies for unknown and standard (differences due to geometry and gamma-ray shielding) and:

(2)

where = decay corrected count rate; = net counts in γ-ray peak corrected for nuclear interferences; λ= desintegration constant; = decay time; = live time of count; = irradiation time; = pulse pileup correction; = live time correction.

Note that the „R-values“ are normally very close to unity, and all units are either SI-based or dimensionless ratios. Since all units are SI, an uncertainty statement can be written down in terms of SI units (or ratios).

All uncertainty components can be completely evaluated as demonstrated by several authors [7,8], among them by Kučera et al. [8]. This is discussed in detail in par. 3.2. This piece of knowledge can be considered as the author´s contribution to the recognition of NAA as a primary method of measurement.

There are no objective criteria how to demonstrate the highest metrological properties of a method. Greenberg suggested [4] that “the highest metrological properties” seems to imply a comparison to other methods of analysis and may include some or all of the following criteria: (1) freedom from bias (or potential accuracy of the method); (2) measurement precision; (3) completeness and magnitude of uncertainty budgets; (4) traceability, and (5) detection limits. In the field of trace element analysis, a comparison between ID-ICPMS and NAA seems appropriate. If we consider the five criteria listed above, the two methods appeared comparable [4]. The direct comparison of the performance of NAA and IDMS at the highest metrological levels has become possible through interlaboratory (key) comparisons conducted by CCQM.  Comparisons of the performance of the two methods in the key comparisons proved similar levels of performance in terms of magnitude of uncertainty statements and agreement (within uncertainties) of the reported values and the accepted values for the studies for a number of elements in various matrices. Hence, in view of similarity of results obtained by NAA and ID-ICPMS, it now seems reasonable to accept that NAA (using the comparator method of standardization) meets all requirements of a primary method of measurement [4].

1.2	Use of NAA in the preparation of reference materials

Reference materials play an essential role in chemical metrology, in laboratory quality assurance procedures. Within this context of quality assurance, reference materials are mostly used for: (1) development, characterization, and validation of analytical methods; (2) internal quality control; (3) external quality assessment.

Reference material (RM) is a material, sufficiently homogeneous and stable with reference to specified properties, which has been established to be fit for its intended use in measurement or in examination of nominal properties. Certified reference material (CRM) is a RM accompanied by documentation issued by an authoritative body and providing one or more specified property values with associated uncertainties and traceabilities, using valid procedures [9]. The property value in chemical analysis is an amount of substance, e.g., an element or a compound.

The preparation of a matrix RM involves the following steps: (1) material selection; (2) collection; (3) drying; (4) crushing and particle size fractionation; (5) homogenization and homogeneity testing; (6) stability testing; (7) and most importantly characterization – certification of analyte contents.

In view of favourable features of NAA, this technique has been extensively used for the most important steps of the preparation of RMs of chemical composition – for determination of element contents in homogeneity and stability testing, and namely for certification of element contents.

In homogeneity and stability testing, the crucial problem is the differentiation of measurement uncertainty from the uncertainty due to sample inhomogeneity and/or instability. Using NAA in its non-destructive mode, so-called instrumental NAA (INAA), this problem may be solved much more easily than in most other analytical techniques. Unlike non-radiometric methods of analysis, uncertainty modeling in NAA is facilitated by the existence of counting statistics, because this parameter is instantly available from each measurement [10].

The task of certification, defined as the assignment of concentration data which approaches as closely as possible the “true value” within uncertainty limits, is one of the most, if not the most demanding challenges. In this task, NAA plays a very important role as demonstrated in Table 1, which gives examples of the use of NAA in the certification of element contents in various environmental and biological reference materials from two world-leading producers, the US National Institute of Standards and Technology (NIST) and the EC Joint Research Centre, the Institute of Reference Materials and Measurements (IRMM, formerly Community Bureau Reference – BCR). The trade mark of a certified reference material (CRM) produced by NIST is “standard reference material” (SRM).

Table 1. Use of NAA in the certification of element contents in various environmental and biological reference materials [3].

| Reference material | Na | Elements certified using NAA | NAA share, %b | Year of production |
| --- | --- | --- | --- | --- |
| NIST SRM 1632 Coal | 14 | As, Cd, Cr, Cu, Hg, Mn, V, Zn | 57 | 1975 |
| NIST SRM 1571 Orchard Leaves | 23 | As, Cd, Cr, Cu, Fe, Hg, K, Mg, Mn, Mo, Na, Ni, P, Rb, Sb, Se, U, Zn | 78 | 1977 |
| NIST SRM 1577 Bovine Liver | 12 | Fe, Hg, K, Mn, Na, Se, Zn | 58 | 1977 |
| NIST SRM 1632b Coal | 24 | Al, As, Ba, Ca, Co, Fe, K, Mg, Mn, Na, Rb, Se, Th, Ti, U, Zn | 67 | 1997 |
| NIST SRM 1633b Coal Fly Ash | 23 | Al, As, Ba, Ca, Cr, Fe, K, Mn, Na, Se, Sr, Th, Ti, U, V | 65 | 1993 |
| NIST SRM 1570a Spinach Leaves | 19 | Al, As, Ca, Co, Cu, Hg, K, Mn, Na, Ni, Se, Sr, Th, V, Zn | 79 | 1996 |
| NIST SRM 1573a Tomato Leaves | 21 | Al, As, Ca, Cd, Co, Cu, Cr, Hg, Fe, K, Mn, Na, Ni, Rb, Sb, Se, V, Zn | 86 | 1995 |
| NIST SRM 1577b Bovine Liver | 18 | Ag, Cd, Cl, Cu, Fe, K, Mg, Mn, Mo, Na, Rb, Se, Zn | 72 | 1991 |
| NIST SRM 1577c Bovine Liver | 20 | Ag, As, Ca, Cd, Co, Cr, Cu, Fe, K, Mg, Mn, Mo, Na, Ni, S, Se, V, Zn | 90 | 2008 |
| BCR CRM 038 Coal Fly Ash | 13 | As, Cd, Co, Cr, Cu, F, Fe, Hg, Mn, Na, Zn | 85 | 1995 |
| BCR CRM 277 Estuarine Sediment | 10 | As, Cr, Hg, Sc, Se, Zn | 50 | 1988 |
| BCR CRM 101 Spruce needles | 9 | Al, Ca, Cl, Mg, Mn, P, Zn | 78 | 1990 |

a - number of elements certified; b – share of NAA in the certification

Table 1 shows that the share of NAA in the certification, which exceeds that of any other analytical technique, does not significantly change with time. Now, after recognition of NAA as a primary method of measurement, it can be expected that the role of NAA in this field of chemical metrology will become even more important.

2	EXPERIMENTAL

The results reported were achieved within last 30 years using facilities and equipment, which underwent many improvements. Therefore, it appears reasonable to list only a range of parameters of facilities available in the author’s laboratory and procedures developed during last 30 years.

Irradiations were carried out in a nuclear reactor at Řež (nowadays LVR-15 reactor), which was operated at 2 – 9 MW power. Thermal, epithermal and fast neutron fluence rates were in the range of 1∙1013 cm-2 s-1 to 8∙1013cm-2 s-1, 5∙1012 cm-2 s-1 to 3∙1013cm-2 s-1 and 7∙1012 cm-2 s-1 to 2∙1013cm-2 s-1, respectively. Short-time irradiation (10 s – 5 min.) was performed in polyethylene transport containers (rabbits) with the aid of a pneumatic transfer facility with a transport time, which was shortened from 25 s to 3.5 s. Long-time irradiation (1 hour to several days) was carried out in Al irradiation cans. For irradiation with epithermal neutrons, special capsules lined with or made of 1 mm-thick cadmium were used.

Gamma-ray spectrometry measurements were carried out with HPGe coaxial detectors with relative efficiency ranging from 11 % to 53 %, FWHM resolution 1.85 – 1.75 keV for the 1332.5 keV photons of 60Co. In the mid-1990’s a well-type HPGe detector and a planar HPGe detector became available. The former detector has an active volume of 150 cm3 and FWHM resolution of 2.1 keV for the 1332.5 keV photons of 60Co, the latter detector has an active area of 200 mm2, thickness 13 mm and FWHM resolution of 510 eV for the 122 keV photons of 57Co. Nowadays these detectors are connected to a Canberra Genie 2000 gamma-spectrometer through the chain of associated linear electronics equipped also with a Canberra 599 Loss Free Counting module. Counting times varied from several seconds to several days as a function of the half-life and activity of radionuclides measured.

The INAA procedures are given in detail in the respective references. Since Kučera and Zeisler [11] demonstrated that RNAA is most frequently the most effective means of optimization in NAA and yields the lowest detection limits and uncertainties, a number of RNAA procedures were also developed and employed. These procedures are briefly described in Tables 2 and 3 and their details can be found in the corresponding references.

Table 2. Single-element RNAA procedures for assay of biological and environmental samples

| Element | Nuclear reaction | Sample decomposition | Separation | Ref. |
| --- | --- | --- | --- | --- |
| Si | 30Si(n,()31Si | Alkaline-oxidative fusion with Na2O2 + NaOH | Distillation of SiF4 | 12 |
| V | 51V(n,()52V | Pre-irradiation dry ashing in air followed by post-irradiation wet ashing in a mixture of H2SO4 + HNO3 + HClO4 | Extraction with N-benzoyl-phenyl-hydroxylamine | 13 |
| Cr | 50Cr(n,()51Cr | Wet ashing in a mixture of HNO3 + HClO4 | Extraction with tribenzylamine | 14 |
| Mn | 55Mn(n,()56Mn | Wet ashing in a mixture of HNO3 + HClO4 | Precipitation of hydrated MnO2 | 15 |
| I | 127I(n,()128I | Alkaline-oxidative fusion with Na2O2 + NaOH | Extraction of elementary I with CHCl3 | 16 |
| Re | 185Re(n,()186Re 187Re(n,()188Re | Microwave assisted wet ashing in a mixture of HNO3 + HF or alkaline-oxidative fusion with Na2O2 + NaOH | Extraction with tetraphenylarsonium chloride or methylethylketone or solid extraction with trioctyl-methylammonium chloride | 17,18 |
| Pt | 198Pt(n,()199Pt→ 199Au | Wet ashing in a mixture of HNO3 + HClO4 | Coprecipitation of elementary Au with Se by ascorbic acid | 19 |
| Tl | 203Tl(n,()204Tl 203Tl(n,2n)202Tl | Wet ashing in a mixture of HNO3 + HClO4 (+ HF) | Extraction with Na diethyldithiocarbamate | 20 |

Table 3. Multi-element RNAA procedures for assay of biological and environmental samples

| Element (nuclide) | Sample decomposition | Separation | Ref. |
| --- | --- | --- | --- |
| Rare earth elements | Alkaline-oxidative fusion with Na2O2 + NaOH | Precipitation with oxalic acid | 21,22 |
| Cu, As, Mo, Cd, Sb | Wet ashing in a mixture of H2SO4 + HNO3 + H2O2 | Extraction with Zn diethyldithiocarbamate | 23,24 |
| Co, Ni (58Co) | Wet ashing in a mixture of HNO3 + HClO4 | Ion exchange chromatography using Dowex 2x8 | 25 |
| Hg, Se | Microwave assisted wet ashing in HNO3 | Extraction with Ni diethyldithiocarbamate and precipitation of elemental Se with ascorbic acid | 23 |
| I, Mn | Alkaline-oxidative fusion with Na2O2 + NaOH | Extraction of elementary I with CHCl3 and precipitation of hydrated MnO2 or Mn extraction with Na diethyldithiocarbamate | 26 |

3	RESULTS AND DISCUSSION

  - 3.1		Self-verification principle of NAA

Among the favourable features of NAA listed in Ch. 1, the possibility to determine a particular element using different neutron induced reactions of its isotopes is a very special one, which has no analogy in other analytical techniques. Since it forms the basis for a unique ability to verify analytical data that NAA produces, Byrne and Kučera termed this special property the self-verification principle [2,27]. Fig. 3 shows the basic elements of the self-verification principle in NAA. The primary source of the independence of the analytical data is the essentially isotopic nature of NAA, i. e., activation of different isotopes of the same element, and the possibility of using different, again essentially independent, isotopic nuclear reactions, such as (n, (), (n,p), (n,(), (n,2n), (n,f) in NAA. For example, it has been shown [2] that more than 25 elements most frequently assayed in biological and environmental matrices can be determined by at least two independent reactions giving gamma emitting radionuclides.

Fig. 3. Block diagram of strategies for self-verification in NAA

Fig. 4 shows an example of self-verification of INAA results for Se in human red blood cells using two independent nuclear reactions 74Se(n,γ)75Se and 76Se(n,γ)77mSe [28].

Fig. 4. Self-verification of INAA results for Se in human red blood cells using two independent nuclear reactions

In addition to the undoubted independence of analytical data provided by NAA using different isotopic reactions, other elements of independence exist in NAA methods. Internally independent routes to analytical information in NAA can be provided by several ways:

i) using the different modes of NAA non-destructive, instrumental NAA (INAA) or radiochemical NAA (RNAA);

ii) using selective activation with thermal, epithermal and fast neutrons (denoted TNAA, ENAA, and FNAA in Fig. 3, respectively);

iii) using selective measurement techniques ((-, (-, and X-ray spectrometry, Compton suppression counting in  (-ray spectrometry, (-( or (-( coincidence counting, ( and Cherenkov counting, and delayed neutron or fission track counting for uranium determination);

iv) using combinations of the above.

Table 4 demonstrates the self-verification principle for results of Tl, U and I using independent routes of their determination by NAA.

Table 4. Results for NIST SRM 1633a Coal Fly Ash using the self-verification principle of NAA [28]

| Element (unit) | Nuclear reaction | NAA mode | This work | NIST value |
| --- | --- | --- | --- | --- |
| Tl (μg g-1) | 203Tl(n,2n)202Tl | RNAAa | 5.70 ( 0.57 | 5.7 ( 0.2 |
|  | 203Tl(n,()204Tl | RNAAb | 5.58 ( 0.25 |  |
|  | 203Tl(n,()204Tl | RNAAc | 5.41 ( 0.67 |  |
|  | 203Tl(n,()204Tl | RNAAd | 5.80 ( 1.14 |  |
| U (μg g-1) | 238U(n,(,(-)239Np | INAA | 10.4 ( 0.3 | 10.2 ( 0.1 |
|  | 238U(n,f)131I | RNAA | 10.3 ( 0.4 |  |
| I (μg g-1) | 127I(n,()128I | INAAe | 0.47 (2) | - |
|  | 127I(n,()128I | RNAAf | 0.441 ( 0.043 |  |
|  | 127I(n,2n)126I | RNAAg | 0.446 ( 0.025 |  |

a: counting of γ-rays of 202Tl; b: counting of Hg X-rays of 204Tl; c:  LSC of 204Tl; d: counting of Cherenkov radiation of 204Tl; e: pre-irradiation separation; f: KOH fusion; g: O2 combustion

Use of the above four strategies (i) – (iv) (and different isotopic routes) provides assurance that the errors to which NAA is liable, like all analytical techniques, will be reduced or at least altered in the alternative route. The more elements of independence are combined, the less the likelihood of repeating  the same error. The great range and variety of possible modes available in NAA should be considered a valuable asset. The belief that multiple data on an element are redundant or methods superfluous is contrary to the principle of consistency through diversity. To conclude, the possibility to determine a particular element using different neutron induced reactions of its isotopes, together with the ability to perform the analysis both by INAA and RNAA, with additional options for internally cross-checking the results obtained by various methodological approaches in NAA enhances the credibility of the data obtained.

3.2	Uncertainty quantification in NAA

Uncertainty (of measurement) is defined [9] as ”non-negative parameter characterizing dispersion of the quantity values being attributed to a measurand, based on the information used”. Its quantification is of utmost importance in all types of measurement, and therefore guidelines with examples were developed for some analytical techniques for better understanding of this evaluation methodology by the analytical community. Recently, the International Atomic Energy Agency (IAEA) took the initiative in developing such guidelines with examples for selected nuclear and nuclear related measurement techniques, involving also NAA, in the form of a technical document (TECDOC). The author participated in the preparation of this TECDOC [29] and also published one of the very first papers on quantification of uncertainty in NAA [30]. For sake of simplicity, only quantification of uncertainty in relative (comparator) NAA will be discussed here.

In general, the sources of standard uncertainty ui can be grouped according to the individual steps of analysis into four categories: (1) preparation of the sample and comparator (standard, neutron fluence rate monitor); (2) irradiation; (3) (-ray spectrometry measurement; and (4) radiochemical separation, if performed. The list of the standard uncertainties of the above categories and their typical values are listed in Table 5.

Table 5. Uncertainty sources and their typical values in NAAa

| Source of uncertainty | Typical ui, %b |
| --- | --- |
| Sample and comparator preparation, u1 |  |
| Mass determination of a sample, u1a | 0.015- 0.1 % |
| Mass determination of comparators, u1b | 0.075 - 0.2 % |
| Mass changes of samples due to moisture uptake during weighing | negligible-1 % |
| Concentration of comparators (standards), purity and stoichiometry of chemicals used for the preparation of standards, u1d | 0.1-several % |
| Variation of isotopic abundance, u1e | negligible for most elements |
| Blank variation and the necessary correction (due to analyte content in an irradiation vial), u1f | 0.1- several % |
| Irradiation, u2 |  |
| Irradiation geometry differences, u2a | < 0.1 - 0.5 % |
| Neutron self-shielding/scattering differences, u2b | ~ 0.1 % in most cases |
| Duration (ttiming) of irradiation, u2c | negligible - 0.3 % |
| Nuclear reaction interferences, u2d | negligible - ~ 1 % |
| Neutron spectrum variations in time and space, u2e | negligible - ~ 1 % |
| Volatilization losses during irradiation, u2f | negligible - several % |
| (-ray spectrometry measurement, u3 |  |
| Counting statistics, u3a | usually 0.2 - 30 % |
| Counting geometry differences, u3b | ~ 0.1 - 5 % |
| Pulse-pileup corrections (random coincidences), u3c | ~ 0.1 - 0.5 % |
| True coincidences (cascade summing), u3d | ~ 1%c |
| Live time corrections, u3e | negligible in most cases |
| - Continuation - |
| Source of uncertainty | Typical ui, %b |
| Decay timing corrections, u3f | negligible in most cases |
| Duration (timing) of counting, u3g | negligible in most cases |
| (-ray self-shielding, u3h | ~ 0.1 - ~ 0.5 % |
| (-ray interferences, u3i | ~ 0.3 - ~ 1 % |
| Peak integration method, u3j | ~ 0.5 - several % (for multiplets) |
| Blank correction (due to counting room/shielding background), u3k | negligible in most cases |
| Radiochemical separation, u4 |  |
| Mass determination of stable carrier and/or radiotracer, u4a | 0.02 - 0.5 % |
| Yield determination, u4b | ~ 0.3 - ~ 0.5 % |
| Isotopic exchange between the radioindicator and radiotracer and/or stable carrier, u4c | negligible in most cases |

a - for details see Refs. 29,30

b - relative standard uncertainty

c - irrelevant in relative standardization

An example of uncertainty evaluation for low-level manganese determination in biological material, which requires the use of RNAA, was reported in [30]. A sample and a Mn standard (comparator) were irradiated simultaneously (with the same neutron fluence), Mn was separated by a radiochemical procedure [15] and counted immediately after separation with a HPGe detector. Afterwards, the manganese standard was counted with the same detector and in the same conditions (counting geometry).

For this case, and considering that the ratios, , and equal to unity, and negligible blank, equations (1) and (2) can be written as

(3)

where  mass fraction of Mn in the sample,  and = net counts in γ-ray peak (corrected for nuclear interferences), and =decay factors (),=sample mass and Y =yield of radiochemical separation.

Since equation (3) involves only multiplication and division of independent quantities, the combined uncertainty uc(cMn) can be calculated according to equation

(4)

Obviously, the quantities given in equation (3) do not directly represent all possible sources of uncertainties as listed in Table 5. Therefore, to be able to demonstrate that the uncertainty budget is complete and corresponds with the measurement equation, it has been suggested to project variables not contained in equation 3 into those that are included in this equation (e.g., all uncertainty sources originating from irradiation and γ-ray spectrometry can be projected into the  and quantities) as schematically depicted in the cause and effect (Ichikawa) diagram in Fig. 5.

Fig. 5. Cause and effect diagram of uncertainty sources in RNAA

Then, sources of uncertainty that are irrelevant in relative standardization or negligible can be left out and the combined uncertainty calculated according to equation (4) as was illustrated in [29,30].

  - 3.3	Use of NAA in the certification of reference materials

Realizing the advantageous features and the indispensable role of NAA in the preparation of reference materials, the author participated in homogeneity testing and certification analyses of numerous environmental and biological RMs of chemical composition using both INAA and RNAA procedures since the late 1980’s. Table 6 shows the main accomplishments that were achieved in

Table 6. Activities in the preparation of reference materials of various producers

| RM code | Producer | Method | Activity | Ref. |
| --- | --- | --- | --- | --- |
| Coal Fly Ash ENO, ECH, EOP, ECO | IRANT, Košice | INAA | Homogeneity tests, certification analyses for up to 32 elements | 31,32 |
| Bovine Liver, Bovine Kidney, Bovine Muscle | CZIM, Bratislava | INAA, RNAA | Homogeneity tests, certification analyses for up to 22 elements, evaluation of certification campaign | 33,34 |
| Green Algae 12-02-02, Lucerne 12-02-03, Wheat Bread Flour 12-02-04, Rye Bread Flour 12-02-05 | IRANT, Košice | INAA, RNAA | Homogeneity tests, certification analyses for up to 27 elements | 35 |
| CRM 7001 Light Sandy Soil, CRM 7002 Light Sandy Soil, CRM 7003 Silty Clay Loam, CRM 7004 Loam | ČMI, Brno | INAA | Homogeneity tests and certification analyses for up to 28 elements, evaluation of certification campaign for total element contents and extractable fractions | 36,37 |
| Apatite Concentrate CTA-AC-1, Fine Fly Ash CTA-FF-1, Oriental Tobacco Leaves CTA-OTL-1, Virginia Tobacco Leaves CTA-VTL-1 | INCT, Warsaw | INAA, RNAA | Certification analyses for up to 34 elements | 38 |
| IAEA A-11 Milk Powder, IAEA H-4 Animal Muscle, IAEA A-13 Freeze Dried Animal Blood, IAEA H-8 Horse Kidney, IAEA H-9 Mixed Human Diet | IAEA, Vienna | RNAA | New data on Cu and Mn | 15,39 |
| Single cell green algae RMs with managed levels of heavy metals IAEA-391, IAEA-392, IAEA-393 | IAEA, Vienna | INAA, RNAA | Certification analyses for up to 22 elements | 40 |
| - continuation - |
| RM code | Producer | Method | Activity | Ref. |
| IAEA Reference Air Filters (RAF)-I, IAEA Reference Air Filters-II | IAEA, Vienna | INAA | Preparation of RAF, homogeneity testing and certification analyses for up to 35 elements, evaluation of the certification campaign | 41-43 |
| IRMM-530 Al-0.1 % Au alloy | IRMM, Geel | INAA | New data on Au content | 44 |
| IRMM-540R, IRMM-541 Uranium-doped glass | IRMM, Geel | INAA | Certification analyses for uranium | 45 |
| NIST SRM-1515 Apple Leaves, NIST SRM-1515 Peach Leaves | US NIST, Gaithersburg | INAA, RNAA | Certification analyses for 28 elements | 46,47 |
| NIST SRM-1573a Tomato Leaves | US NIST, Gaithersburg | INAA, RNAA | Certification analyses for 28 elements | 38,47 |
| NIST SRM-1570a Spinach Leaves | US NIST, Gaithersburg | INAA, RNAA | Certification analyses for 20 elements | 47 |
| NIST SRM-2783 Air Particulate on Filter Media | US NIST, Gaithersburg | INAA | Certification analyses for 23 elements | 47 |
| NIST SRM-1648 Urban Particulate Matter | US NIST, Gaithersburg | INAA | New data for V and Mn | 47,48 |
| NIST SRM-1577c Bovine Liver | US NIST, Gaithersburg | RNAA | Certification analyses for V and Ni | 49 |

co-operation with Czechoslovak (later Czech) producers of RMs and those from abroad, namely the Polish Institute of Nuclear Chemistry and Technology (INCT), the International Atomic Energy Agency (IAEA), the Institute of Reference Materials and Measurement (IRMM – Joint Research Centre of the European Commission), and the US National Institute of Standards and Technology (NIST). Only those results are included in Table 6 that were achieved during the preparation of the individual RMs and within their certification campaign, i.e., prior to certified and/or noncertified values of element contents were divulged. Exceptions are analyses that were carried out to confirm and/or to correct the existing certified values (Cf. par. 3.4). It may also be mentioned that only those results are shown in Table 6 that were published in a refereed journal or were acknowledged in the respective certificates (except for participations in the certification of IRMM RMs). Since it is not possible to present all author’s results in this work in detail, only selected examples are given to illustrate and/or to highlight some of the major achievements.

Regarding testing homogeneity of element contents, Fig. 6  demonstrates a wide dynamic range of INAA, which allows to test inhomogeneities of a number of elements, whose contents differ from per cent down to mg·kg-1, as was achieved in homogeneity testing of soil CRM No. 7004 [36].

Fig. 6. Relative inhomogeneity, sh, of elements with various contents in soil CRM No. 7004 as determined by INAA

Table 7 gives an example of multielement performance of both INAA and RNAA procedures applied in the author’s laboratory in the certification campaign of NIST SRM-1573a Tomato Leaves.

Table 7. Element contents in NIST SRM 1573a Tomato Leaves (dry mass)

| Element, | This work (N=6) | NIST valueb |
| --- | --- | --- |
| unit | Method | Mean Ua |  |
| As, ng·g-1 | RNAA | 109.6 4.1 | 112 4 |
| Ba, g·g-1 | INAA, 139Ba | 57.1 12.2 | (63) |
|  | INAA, 131Ba | 63.0 3.8 |  |
| Br, g·g-1 | INAA | 1189 33 | (1300) |
| Ca, % | INAA | 4.970 0.143 | 5.05 0.09 |
| Cd, g·g-1 | RNAA | 1.55 0.06 | 1.52 0.04 |
| Cl, g·g-1 | INAA | 6647 102 | - |
| Co, g·g-1 | INAA | 0.561 0.021 | 0.57 0.02 |
| Cr, g·g-1 | INAA | 1.95 0.03 | 1.99 0.06 |
| Cs, ng·g-1 | INAA | 51.6 5.1 | (53) |
| Cu, g·g-1 | RNAA | 4.64 0.14 | 4.70 0.14 |
| Eu, ng·g-1 | INAA | 36 3 | - |
| Fe, g·g-1 | INAA | 369 17 | 368 7 |
| Hf, ng·g-1 | INAA | 128 4 | (140) |
| Hg, ng·g-1 | RNAA | 33.4 1.2 | 34 4 |
| K, % | INAA | 2.681 0.064 | 2.70 0.05 |
| La, g·g-1 | INAA | 2.26 0.06 | (2.3) |
| Mn, g·g-1 | INAA | 247.1 4.7 | 246 8 |
| Mo, ng·g-1 | RNAA | 463 33 | (460) |
| Na, g·g-1 | INAA | 133 4 | 136 4 |
| Rb, g·g-1 | INAA | 14.6 0.3 | 14.89 0.27 |
| Sb, ng·g-1 | RNAA, 122Sb | 61.9 8.8 | 63 6 |
|  | INAA, 124Sb | 59.9 9.7 |  |
| Sc, ng·g-1 | INAA | 101 3 | (100) |
| Se, ng·g-1 | RNAA | 57.6 5.8 | 54 3 |
| Sm, ng·g-1 | INAA | 0.19 0.02 | (0.19) |
| Sr, g·g-1 | INAA | 80.4 6.9 | (85) |
| Th, ng·g-1 | INAA | 116 6 | (120) |
| V, g·g-1 | INAA | 0.81 0.20 | 0.835 0.01 |
|  | RNAA | 0.78 0.06 |  |
| Zn, g·g-1 | INAA | 30.1 1.1 | 30.9 0.7 |

a - expanded uncertainty (coverage factor k=2)

b - values with uncertainty are certified, noncertified values are in parenthesis

An excellent agreement can be seen in Table 7 between the author’s results and NIST certified or noncertified values of element contents. This proves that highly accurate INAA and RNAA procedures were developed in the author’s laboratory. The self-verification principle of NAA was employed whenever possible (Cf. values for Ba and Sb in Table 7).

Another example of multielemental capability of NAA and the possibility to perform analysis non-destructively concerns certification analyses of NIST SRM-2783 Air Particulate on Filter Media as illustrated in Fig. 7.

Fig. 7. Comparison of NIST certified values (diamonds, ref. 47) with INAA results achieved in the author’s laboratory for loaded filters (squares) of NIST SRM 2783

It is obvious that the author’s results agree with the NIST certified values within uncertainty margins.

Sometimes it is not the multielemental capability of NAA, but the requirement of accurate determination of selected elements, which is needed in the certification of specific RMs. For instance, when IRMM prepared two RMs of uranium doped glass for fission track dating, IRMM-540R and IRMM-541, several laboratories experienced in uranium determination using various analytical techniques were needed. One of them (coded Lab 07 in Table 8) was the author’s laboratory using INAA. Table 8 shows results of characterization tests obtained at 8 laboratories.

Table 8. Results of characterization tests obtained at 8 laboratories on IRMM-540R samples (top) and IRMM-541 samples (bottom) [45]

Recently, NIST has prepared a freeze-dried powdered bovine liver tissue SRM to renew supply of the benchmark material SRM 1577b Bovine Liver. The collection and preparation of this material was performed under strict control of contamination resulting in a unique SRM with very low levels of some trace elements typical for uncontaminated biological tissues. The inclusion of RNAA procedures for the determination of V and Ni at the author’s laboratory was essential to meet NIST criteria [50] for assignment of mass fractions of these elements. Table 9 compares RNAA results for V and Ni with the NIST certified values, which were obtained by combining the values found in this work [49] with ICP-MS and preseparation-NAA values obtained previously at NIST.

Table 9. Comparison of RNAA results for V and Ni with NIST certified values for NIST SRM 1577c Bovine Liver.

| Element, ng·g-1 | V | Ni |
| --- | --- | --- |
| Author’s resultsa | 8.4 ± 1.0 | 44.0 ± 6.4 |
| NIST valuea | 8.17 ± 0.66 | 44.5 ± 9.2 |

a – expanded uncertainties with coverage factor of 2 are given

3.4	Use of NAA in the detection of errors

It is not only the frequent use of NAA that reflects its importance in the RM certification, but the technique has many times helped in resolving of discrepant analytical data and there are also several examples that the consistent application of NAA advantages lead to the corrections of originally certified values. Recommended values of Cu and Mn in RM Milk Powder A-11 prepared by the International Atomic Energy Agency (IAEA) at the end of the 1980´s from an open intercomparison were questioned, because there was disagreement between NAA and AAS results [39]. A follow-up study was organized later in selected laboratories, largely using NAA. The NAA author’s results also helped in establishing newly assigned values of Cu and Mn in IAEA RM A-11 Milk Powder [51] as demonstrated in Table 10.

Table 10. Comparison of old and new values for Cu and Mn in IAEA RM A-11 Milk Powder

| Element, unit | Originally assigned value [51] | Author’s results [15,39] | Newly assigned value [51] |
| --- | --- | --- | --- |
| Cu, ng·g-1 | 838 ± 450 | 400 ± 20 | 378 ± 31 |
| Mn, ng·g-1 | 377 ± 177 | 258 ± 14 | 257 ± 6 |

While the above case concerned the correction of data obtained in an open intercomparison, the author also noticed differences of his INAA results and NIST values for Mn and V in NIST SRM 1648 Urban Particulate Matter and succeeded in demonstrating the accuracy of his data. In 1998, his finding has led to a change of the NIST values assigned in 1982 to the new values as given in Table 11.

Table 11. Comparison of old and new values for Mn and V in NIST SRM-1648 Urban Particulate Matter

| Element, unit | Original NIST valuesa [47 ] | Author’s results [48 ] | New NIST valuesa [47 ] |
| --- | --- | --- | --- |
| Mn, µg·g-1 | (860) | 768 ± 18 | 786 ± 17 |
| V, µg·g-1 | 140 ± 3 | 128 ± 4 | 127 ± 7 |

a-NIST values with uncertainty are certified, noncertified values are in parenthesis

Noteworthy, this was only the second case, where data for environmental and/or biological reference materials obtained outside NIST has led to a correction of the NIST certified values of element contents.

Al-0.1% Au alloys with known uncertainty of their Au content are widely used in the k0-NAA standardization and other nuclear-analytical applications for neutron spectrum and fluence measurements. Inconsistency of the certified Au content in the IRMM-530 RM with declared Au value in the 602D Al-0.1% Au alloy produced by Reactor Experiments (R/X, USA) was found in several laboratories, indicating that one of the values was biased by about 3-5 %. This would have a dramatic effect on the accuracy of neutron flux measurements, and consequently on the accuracy of  k0-NAA results. Therefore, the Au content and its homogeneity in the alloys were examined in the author’s laboratory. Table 12 summarizes the results obtained.

Table 12. Results of Au determination in Al-0.1 % Au alloys

| Material | Overall mean ± uc | N | Catalogue value |
| --- | --- | --- | --- |
| IRMM-530 foil | 0.0999 ± 0.00038 | 19 | 0.1000 ± 0.0020 |
| IRMM-530 wire | 0.1003 ± 0.00036 | 34 | 0.1000 ± 0.0020 |
| R/X 602D wire | 0.1073 ± 0.00065 | 24 | 0.1124 |

The mean value found in the author’s laboratory for the IRMM-530 foil and wire was in excellent agreement with the IRMM certified value, while that for the R/X wire was found by 4.5 % lower compared to the declared value.

4 CONCLUSIONS AND OUTLINE OF FUTURE RESEARCH AND EDUCATION ACTIVITIES

The prominent position of NAA in chemical metrology, namely in homogeneity tests and certification of element contents in reference materials, use of the technique in the detection of errors, etc., has been demonstrated. It can be expected that after the recent recognition of NAA as a primary method of measurement, the important and indispensable position of NAA among other trace element analytical techniques in the field of chemical metrology will further be enhanced.

Until now, however, only non-destructive, instrumental NAA (INAA) with relative (comparator) standardization has been given the status of primary method of measurement. Therefore, it is desirable to achieve the same status for other modes of NAA, such as NAA with radiochemical separation (RNAA), NAA with k0 standardization, etc. Theoretically, this seems easy for RNAA, because equation (1) can be changed as

(4)

where = yield of chemical separation, which can be determined in several ways and a complete uncertainty budget can be provided for determination of . However, it necessary to convince CCQM again that for RNAA: (1) a complete uncertainty statement can be written down in terms of SI units (or ratios) and (2) the method has the highest metrological properties. For this reason, it appears necessary to continue in maintaining and developing of such RNAA procedures in which the chemical yield of separation can be determined directly for each measurement, i.e., by using a suitable radiotracer, reactivation of the added inactive carrier or with the aid of another primary method of measurement, e.g., by gravimetry as frequently exemplified in the author’s laboratory. From the metrological point of view, determination of the chemical yield of separation in “model” experiments appears rather problematic.

For k0-NAA standardization, which is not discussed in this work, it will be somewhat more complicated to prove that this standardization mode meets all requirements of a primary method of measurement as given in par. 1.1, but theoretically it is feasible. Implementation and validation of this standardization method in NAA has recently been started in the author’s laboratory.

In education, the goals given in the above motto can be accomplished in two stages. First is the basic, theoretical education concerning principles of analytical methods, and in the context of this work, especially nuclear analytical methods. This is well cover by several courses pursued at the Faculty of Nuclear Sciences and Physical Engineering of CTU in Prague, which can generate analysts “with a good insight into the merits and demerits of a given analytical technique”. However, “the knowledgeable analyst” from the above motto also means an analyst with some, as much as possible, experience. The required experience can only be obtained in practical exercises. This can be best achieved by inclusion of students into research projects and in postgraduate studies, preferably in co-operation with well-experienced tutors. Such activities should start as early as possible in the curricula, because they represent one of the most efficient instruments in motivating the students and getting them acquainted with the current state-of-art of the individual disciplines.

In the field of chemical metrology, namely in quality assurance of nuclear and other analytical methods, building and increasing awareness of this rapidly developing scientific field appears a very important, if not even imperative task. This topic is only partly covered in the existing courses at the Faculty of Nuclear Sciences and Physical Engineering of CTU in Prague, including that one of the author for postgraduate students. For instance, the concept of uncertainty and traceability of chemical measurements (analyses), which form the basis of comparability of these measurements among laboratories worldwide needs more attention. The author is prepared to develop such a course, as has also been discussed with head of the Nuclear Chemistry Department of the Faculty of Nuclear Sciences and Physical Engineering of CTU in Prague. It can be expected that such a course may be attractive not only for students of this department.

Acknowledgements

Several sources of support are highly acknowledged. They include:

- co-workers from my present and former affiliations, colleagues from domestic institutions and those from abroad

- several of my diploma-work and Ph.D. students, who have done some experiments and contributed to their evaluation

- local grant agencies and those from abroad, which providing the financial support of my research work

- the Faculty of Nuclear Sciences and Physical Engineering of CTU in Prague for giving me an opportunity to present results of my research to its students and to attract them to get involved in research conducted in my laboratory.

References

- BYRNE, A.R.: Review of neutron activation analysis in the standardization and study of reference materials including its application to radionuclide reference materials. Fresenius J. Anal. Chem., 345, 1993, pp. 144-151.

- BYRNE, A.R.- KUČERA, J.: Role of the self-validation principle of NAA in quality assurance of bioenvironmental studies and in the certification of reference materials, Proc. Int. Symp. on Harmonization of Health Related Environmental Measurements Using Nuclear and Isotopic Techniques, Hyderabad, India, 4-7 November 1996, IAEA Vienna, 1997, paper IAEA-SM-344/8, pp. 223-238.

- IHNAT, M.- BYRNE, A.R.- KUČERA, J.- QUEVAUVILLER, P.: Certifi-cation of elements. In: Reference Materials for Chemical Analysis. Certification, Availability and Proper Usage (Eds. STOEPPLER, M.- WOLF, W.R.- JENKS, P.J.), Wiley-VCH Verlag GmbH, Weiheim, New York, Chichester, Brisbane, Singapore, Toronto, 2001, pp. 60-75.

- GREENBERG, R. R.: Pushing the limits of NAA: Accuracy, uncertainty and detection limits. J. Radioanal. Nucl. Chem., 278, 2008, pp. 231-240.

- Report of the 4th Meeting (February 1998) of CCQM, BIPM, Sevres, France (February 1999), ISBN 92-822-2164-4, p. 71.

- QUINN, T.J.: News from the BIPM. Metrologia, 36, 1999, pp. 65-69.

- GREENBERG, R.R.- LINDSTROM, R.M.- SIMONS, D.S.: Instrumental neutron activation analysis for certification of ion-implanted arsenic in silicon. J. Radioanal. Nucl. Chem., 245, 2000, pp. 57-63.

- KUČERA, J.- BODE, P.- ŠTĚPÁNEK, V.: The 1993 ISO Guide to the expression of uncertainty in measurement applied to NAA. J. Radioanal. Nucl. Chem., 245, 2000, pp. 115-122.

- International Vocabulary of Metrology – Basic and General Concepts and Associated Terms (VIM), 3rd edition, ISO/IEC Guide 99:2007, Geneva.

- BYRNE, A.R.- HEINZ-GROBECKER, K.- ISENGARD, H.-D.- KUČERA, J.- PAUWELS, J.- RÜCKHOLD, S.- STOEPPLER, M: Control of Material Properties. In: Reference Materials for Chemical Analysis. Certification, Availability and Proper Usage (Eds. STOEPPLER, M.- WOLF, W.R.- JENKS, P.J.), Wiley-VCH Verlad GmbH, Weiheim, New York, Chichester, Brisbane, Singapore, Toronto, 2001, pp. 31-48.

- KUČERA, J.- ZEISLER, R.: Do we need radiochemical separation in activation analysis? J. Radioanal. Nucl. Chem., 262, 2004, pp. 255 - 60.

- KUČERA, J.- ZEISLER, R.: Low-level determination of silicon in biological materials using radiochemical neutron activation analysis, J. Radioanal. Nucl. Chem., 263, 2005, pp. 811-816.

- BYRNE, A.R.- KUČERA, J.: Radiochemical neutron activation analysis of traces of vanadium in biological samples: A comparison of prior dry ashing with post-irradiation wet ashing, Fresenius J. Anal. Chem., 340, 1991, pp. 48 - 52.

- KUČERA, J.- BENCKO, V.- TEJRAL, J.- BORSKÁ, L.- SOUKAL, L.- ŘANDA, Z.: Biomonitoring of occupational exposure: Neutron activation determination of selected metals in the body tissues and fluids of workers manufacturing stainless steel vessels , J. Radioanal. Nucl. Chem., 259, 2004, pp. 7 - 11.

- KUČERA, J.-  SOUKAL, L.- FALTEJSEK, J.: Low level determination of manganese in biological reference materials by neutron activation analysis, J. Radioanal. Nucl. Chem., 107, 1986, pp. 361- 369.

- KUČERA, J.- IYENGAR, G.V.- ŘANDA, Z.- PARR, R.M.: Determination of iodine in Asian diet samples by epithermal and radiochemical neutron activation analysis, J. Radioanal. Nucl. Chem., 259, 2004, pp. 505 - 509.

- KUČERA, J.- BYRNE, A.R.- MIZERA, J.- LUČANÍKOVÁ, M.- ŘANDA, Z.: Development of a radiochemical neutron activation analysis procedure for determination of rhenium in biological and environmental samples, J. Radiaoanal. Nucl. Chem., 269, 2006, pp. 251-257.

- MIZERA, J.- KUČERA, J.- ŘANDA, Z.- LUČANÍKOVÁ, M.: Advanced liquid and solid extraction procedures for ultratrace determination of rhenium by radiochemical neutron activation analysis, Czech J. Phys., 56, 2006, Suppl. D, pp. D315-D321.

- KUČERA, J.- DROBNÍK, J.: Determination of platinum in urine and serum after the administration of cisplatin by neutron activation analysis, J. Radioanal. Chem., 75, 1982, pp. 71 – 80.

- KUČERA, J.- VOBECKÝ, M.- SOUKAL, L.- ZÁKOUCKÝ, D.- VÉNOS, D.: Low level determination of thallium in biological and environmental reference materials by RNAA using several counting methods, J. Radioanal. Nucl. Chem., 217, 1997, pp. 131 - 137.

- KUČERA, J. - MIZERA, J.- ŘANDA, Z.- VÁVROVÁ, M.: Pollution of agricultural crops by the lanthanides, thorium and uranium studied by instrumental and radiochemical neutron activation analysis, J. Radioanal. Nucl. Chem., 271, 2007, pp. 581-587.

- ŘANDA, Z.- KUČERA, J.- SOUKAL, L.: Elemental characterization of the new Czech meteorite „Morávka“ by neutron and photon activation analysis, J. Radioanal. Nucl. Chem., 257, 2003, pp. 275 - 283.

- KUČERA, J.- SOUKAL, L.: Determination of As, Cd, Cu, Hg, Mo, Sb, and Se in biological reference materials by radiochemical neutron activation analysis, J. Radioanal. Nucl. Chem., 168, 1993, pp. 185 - 199.

- LUČANÍKOVÁ, M.- KUČERA, J.- ŠEBESTA, F.- JOHN, J.: Use of new composite materials for separation in RNAA, J. Radioanal. Nucl. Chem., 269, 2006, pp. 463-468.

- KUČERA, J.-  BYRNE, A.R.: Nickel determination in biological materials at ultratrace level by fast neutron radiochemical activation analysis, J.  Radioanal. Nucl. Chem., 168, 1993, pp. 201 - 213.

- KUČERA, J.- MIZERA, J.- REPINC, U.- SMODIŠ, B.: Simultaneous low-level determination of iodine and manganese by radiochemical neutron activation analysis, Czech. J. Phys., 56, 2006, Suppl. D, pp. D151-D157.

- KUČERA, J.: A new role for neutron activation analysis in harmonization and improving quality of chemical analysis. Acta Polytechnica (Prague), 35, 1995, pp. 103-104.

- KUČERA, J.: Methodological developments and applications of neutron activation analysis, J. Radioanal. Nucl. Chem., 273, 2007, pp. 273-280.

- KUČERA, J.- BODE, P.- ŠTĚPÁNEK, V.: Uncertainty evaluation in instrumental and radiochemical neutron activation analysis. In: Quantifying uncertainty in nuclear analytical measurements, IAEA-TECDOC 1401, IAEA Vienna, July 2004.

- KUČERA, J.- BODE, P.- ŠTĚPÁNEK, V.: The 1993 ISO Guide to the expression of uncertainty in measurement applied to neutron activation analysis, J. Radioanal. Nucl. Chem., 245, 2000, pp. 115 - 122.

- KUČERA, J.- SOUKAL, L.: Homogeneity tests and certification analyses of coal fly ash reference materials by INAA, J. Radioanal. Nucl. Chem., 121, 1988, pp. 245-259.

- KUČERA, J.- SOUKAL, L.: Homogeneity tests and certification analyses of the IRANT coal fly ash reference material ECO by instrumental neutron activation analysis, J. Radioanal. Nucl. Chem., 134, 1989, pp. 209 - 219.

- KUČERA, J.- MADER, P.- MIHOLOVÁ, D.- CIBULKA, J.- POLÁKOVÁ, M.- KORDÍK, D.: Preparation of the Bovine Liver 12-02-01 reference material and the certification of element contents from an interlaboratory comparison, Fresenius J. Anal. Chem., 338, 1990, pp. 66 - 71.

- KUČERA, J.- MADER, P.- MIHOLOVÁ, D.- CIBULKA, J.- FALTEJSEK, J.- KORDÍK, D.: Preparation of the bovine kidney and bovine muscle reference materials and the certification of element contents from interlaboratory comparisons, Fresenius J. Anal. Chem., 352, 1995, pp. 66-72.

- KUČERA, J.- SOUKAL, L.: Neutron  activation of new botanical reference materials, Part  II - Evaluation of Czechoslovak RMs Green Algae, Lucerne, Wheat and Rye Bread  Flour,  Fresenius J. Anal. Chem., 345, 1993, pp. 193 - 197.

- KUČERA, J.- SYCHRA, V.- HORÁKOVÁ, J.- SOUKAL, L.: Use of INAA in the preparation of a set of soil reference materials with certified values of total element contents, J. Radioanal. Nucl. Chem., 215, 1997, pp. 147 - 155.

- KUČERA, J.- SYCHRA, V.- KOUBEK, J.: A set of four soil reference materials with certified values of total element contents and their extractable fractions, Fresenius J. Anal. Chem., 360, 1998, pp. 402 - 405.

- KUČERA, J.: Elemental characterization of new Polish and U.S. NIST geological, environmental and biological reference materials by neutron activation analysis and comments on the  methodology of interlaboratory comparisons, Chem. Anal. (Warsaw), 40, 1995, pp. 405 - 421.

- DE GOEIJ, J.J.M.- KOSTA, L.- BYRNE, A.R.- KUČERA, J.: Problems in current  procedures for establishing recommended  values of  trace-element levels in biological reference  materials, illustrated by IAEA Milk Powder A-11, Anal. Chim. Acta, 146, 1983, pp. 161 - 169.

- ZEISLER, R.- DECKNER, R.- ZEILLER, E.- DOUCHA, J.- MADER, P.- KUČERA, J.: Single cell green algae reference materials with managed levels of heavy metals, Fresenius J. Anal. Chem., 360, 1998, pp. 429 - 432.

- KUČERA, J.- PARR, R.M.- SMODIŠ, B.- FAJGELJ, A.- MATTIUZZI, M.- HAVRÁNEK, V.: Use of INAA, PIXE and XRF in homogeneity testing of new IAEA reference air filters, J. Radioanal. Nucl. Chem., 244, 2000, pp. 121 - 126.

- KUČERA, J.- SMODIŠ, B.- BURNS, K.- DE REGGE, P.- CAMPBELL, M.- HAVRÁNEK, V.- MAKAREWICZ, M.- TOERVENYI, A.- ZEILLER, E.: Preparation and characterization of a set of IAEA reference air filters for quality control in air-pollution studies, Fresenius J. Anal. Chem., 370, 2001, pp. 229 - 233.

- KUČERA, J.- HAVRÁNEK, V.- KRAUSOVÁ, I.: Preparation and characterization of a new set of IAEA reference air filters using INAA, PIXE and RBS, J. Radioanal. Nucl. Chem., in press (2009).

- KUČERA, J.- SIMONITS, A.- DE CORTE, F.- BELLEMANS, F.: High accuracy determination of Au by NAA in Al-0.1 % Au alloys, Proc. 2nd Int. k0 Users Workshop, September 30-October 3, 1996, Ljubljana, Slovenia,  pp. 15-17.

- ROEBBEN, G.: The certification of the uranium-doped oxide glasses IRMM-540R and IRMM-541, IRMM Internal Report, Geel, Belgium, 2006.

- KUČERA, J.- SOUKAL, L.- HORÁKOVÁ, J.: Neutron activation of new botanical reference materials, Part I - U.S.NIST Apple and Peach Leaves SRMs, Fresenius J. Anal. Chem., 345, 1993, pp. 188 - 192.

- Certificate of Analysis; available at URL: http://ts.nist.gov/measurementservices/referencematerials/index.cfm

- KUČERA, J.- SOUKAL, L.: Low uncertainty determination of manganese and vanadium in biological and environmental reference materials, Fresenius J. Anal. Chem., 360, 1998, pp. 415 - 418.

- ZEISLER, R.- TOMLIN, B. E.- MURPHY, K. E.- KUČERA, J.: Neutron activation analysis with pre- and post-irradiation chemical separation for the value assignments of Al, V, and Ni in the new NIST SRM 1577c, J. Radioanal. Nucl. Chem., in press (2009).

- MAY, W.E.-GILLS, T.E.-PARRIS, R., et al.: Definitions of Terms and Modes Used at NIST for Value-Assignment of Reference Materials for Chemical Measurements, NIST Special Publication 260-136, US NIST, Gaithersburg, 1999.

- BYRNE, A. R.- CAMARA-RICA, C.- CORNELIS, R., et al.: Results of a co-ordinated research programme to improve the certification of IAEA milk powder A-11 and animal muscle H-4 for eleven “difficult” trace elements, Fresenius Z. Anal. Chem.,  326, 1987, pp. 723-729.

- IYENGAR, V.- WOITTIEZ, J.: Trace Elements in Human Clinical Specimens: Evaluation of Literature Data to Identify Reference Values, Clin. Chem., 34, 1988, pp. 474-481.

Curriculum vitae

doc. Ing. Jan Kučera, CSc.

Personal:

Born on 4th March 1946 in Prague, Czechoslovakia

Education and degrees:

2003 – Associate Professor (doc.), Faculty of Nuclear Sciences and Physical Engineering (FNSPE) of the Czech Technical University in Prague (CTU)

1979 – Ph.D. (CSc.), FNSPE, CTU

1969 – M.Sc. (Ing.), FNSPE, CTU

Professional experience:

1994 – now:		Nuclear Physics Institute, ASCR, Řež, NAA Group Leader

2003 – now:	Centre for Radiochemistry and Radiation Chemistry of FNSPE, CTU, senior research worker (part-time)

1993 – 1994:	Czech Environmental Institute, Prague, Head of NAA Laboratory in Řež

1973 – 1993:	Nuclear Research Institute, Řež, scientific fellow, scientific worker, senior scientific worker, Head of NAA Department since 1977

1971 – 1973:	Geological Survey of the Czechoslovak Uranium Industry, Příbram, research worker

1969 – 1970:	FNSPE, CTU, assistant lecturer

Research areas:

- Instrumental and radiochemical neutron activation analysis

- Instrumental and radiochemical photon activation analysis

- Radiochemical separations

- Application of the above techniques in environmental control, biomedicine geo- and cosmochemistry, and chemical metrology, namely in the preparation of reference materials of chemical composition and in quality assurance

Main projects within last 5 years:

- „Classical and holistic k0 standardization in neutron activation analysis“, grant 202/09/0363 provided by the Czech Science Foundation (CSF), 2009-2011, chief scientific investigator

- IUPAC project „International harmonized protocol for standard preparation, irradiation and measurement for assuring metrologically traceable results in neutron activation analysis”, 2008-2009, chief scientific investigator for the Czech Republic

- „Use of nuclear analytical methods in studying the 99Tc and Re interaction in the environment“, grant 203/08/1276 provided by CSF, 2008-2010, chief scientific investigator

- “Neutron activation analysis of short-lived radionuclides for trace element determination”, bilateral project with the J. Stefan Institute, Ljubljana, Slovenia, 2005-2006 (chief scientific investigator for the Czech Republic

- “Use of solid extraction in radiochemical neutron activation analysis”, grant 203/04/0943 provided by CSF, 2004-2006, co-investigator

- “Virtual Institute for Reference Materials”, contract within the EC 5th Framework – Growth, 2003-2005, Czech National Contact Point

- “Determination of 129I in the environment and wastes of nuclear power plants by neutron activation analysis”, grant A4048301 provided by the Grant Agency of the Academy of Sciences of the Czech Republic, 2003-2005, chief scientific investigator

Education experience:

- Seminar for graduate level “Neutron activation analysis of biological and environmental samples” at FNSPE, CTU since 1988

- Course for postgraduate level “Nuclear analytical methods for environmental control and monitoring” at FNSPE, CUT since 1996

- Lectures on neutron activation analysis, gamma-spectrometry and environmental monitoring, and further to assist laboratories in analytical quality assurance and laboratory accreditation within 14 expert missions (1 week to 1 month) for the International Atomic Energy Agency (IAEA) in 10 countries

- Supervising students´ research projects (two defended), master theses (three defended) and Ph.D. theses (one defended, two running)

- Supervising three IAEA fellows for 6 months and one IAEA fellow for 3 months

Other professional activities:

International: - Member of the Editorial Board of the Journal of Radioanalytical and Nuclear Chemistry

- Member of the International Committee of the series of conferences Nuclear Analytical Methods in the Life Sciences (NAMLS-IC) since 1993, chairman of NAMLS-IC 2002-2008

- Member of the International Committee on Activation Analysis of the series of conferences Modern Trends in Activation Analysis (ICAA/MTAA) since 1996

- Member of the International Committee of the series of International ko-Users Workshops since 2001

National: - Member of the Boards for the State examinations and for defence of Ph.D. theses in Nuclear Chemistry and Analytical Chemistry at FNSPE, CTU and member of the Board for the State examinations and defence of master theses in Nuclear Chemical Engineering at FNSPE, CTU

- Member of the Scientific Board of the Nuclear Physics Institute ASCR, Řež, 2001-2003

- Chairman of the Scientific Board of the Nuclear Physics Institute, ASCR, Řež, 2004-2007

- Chairman of the Board of the Nuclear Physics Institute ASCR, public research institution, since 2008

- Member of the Academic Council of the Academy of Sciences of the Czech Republic since 2004

- Member of the following Learned societies:

- Ioannes Marcus Marci (Czech) Spectroscopic Society since 1975, Head, Task Group for Reference Materials and Standards since 1989

- American Society for Testing and Materials - Task Group on Nuclear Methods of Chemical Analysis since 1989

- EURACHEM – CZ since 2001

- Conference organizing:

- Chairman of the Organizing Committee of the 1st Czechoslovak Conference on Reference Materials, Buřany, 1990

- Chairman of the Organizing Committee of the 5th Int. Conf. on Nuclear Analytical Methods in the Life Sciences (NAMLS-5), Prague, 1993

- Co-chairman of the Organizing Committee of the 13th-15th Int. Radiochemical Conferences, Mariánské Lázně in 1998, 2002, 2006

- Co-chairman of the Organizing Committee of national conferences on Reference Materials and Proficiency Testing in Chemical Analysis II and III, Medlov, in 2005 and 2008

Publications:

Author or co-author of 1 international monograph, 6 chapters in international monographs, 1 chapter in a Czech monograph, 126 papers in refereed international journals, 27 papers in Czech refereed journals, 49 contributions in proceedings of international conferences, 19 contributions in proceedings of Czech conferences, 43 research reports, 4 specialized courses textbooks. The database Web of Knowledge lists 504 citations of the above papers as of March 2009.

Co-editor of 1 technical document (TECDOC) of the International Atomic Energy Agency, 4 Proceedings of international conferences published in refereed journals, 4 Booklets of Abstracts of international conferences.

Honours and awards:

2002 – Ioannes Marcus Marci Medal for outstanding achievements in the field of nuclear analytical methods

2002 – Silver Medal of FNSPE, CTU for pedagogical activities

2006 – Hevesy Medal Award for outstanding contributions to the field of radioanalytical and nuclear chemistry

# Text boxes

CIPM Consultative Committees

for Electricity and Magnetism (CCEM)

for Photometry and Radiometry (CCPR)

for Thermometry (CCT)

for Length (CCL)

for Time and Frequency (CCTF)

for Ionizing Radiation (CCRI)

for Units (CCU)

for Mass and Related Quantities (CCM)

for Amount of Substance (CCQM)

- Established in 1993

for Acoustics, Ultrasound & Vibration (CCAUV)

Motto:

„THE ANALYST is the most important component of any analytical process. With a good insight into the merits and demerits of a given analytical technique and available instrumentation, only a knowledgeable analyst may be expected to generate reliable analytical results.“ [52]